DENTIA

Imaging Capture, Groq AI Analysis & Automated Chart Integration

Technical Architecture & Hardware Report

Prepared for: Dentia Clinical & System Manual

Scope: X-ray sensors, panoramic/OPG machines, intraoral cameras & webcam video — USB capture, Groq AI vision analysis, AI-generated clinical notes, and automated odontogram chart updates.

Regional focus: Pakistan — commonly available imaging hardware

# 1. Executive Summary

This report defines how Dentia captures clinical images and video directly from chairside hardware — intraoral X-ray sensors, panoramic (OPG) machines, and intraoral or webcam cameras — and routes that media into the “Imaging & X-Rays” module without manual file hunting. From there, a Groq-hosted vision-language model observes each image or key video frame, produces a structured clinical read, drafts an AI clinical note, and proposes chart updates that the dentist reviews and confirms before anything is saved to the patient record.

The design keeps three concerns separate: (1) how a physical device gets its image onto the computer running Dentia, (2) how Dentia's backend hands that image to Groq for analysis, and (3) how the AI's output is safely converted into SOAP notes and odontogram changes. Each is covered below, along with a shortlist of imaging hardware commonly available in Pakistan that fits this workflow.

# 2. Objective & Scope

This document answers four practical questions the clinical and engineering teams need before building the capture pipeline:

- How does an X-ray sensor, scanner, or camera plugged into a USB port get its output directly into Dentia's Imaging & X-Rays module — without the dentist manually saving and re-uploading a file?

- How does a live video feed (intraoral camera or webcam) get captured and handed to the AI for observation?

- How does Groq analyze that image or video and turn it into an AI clinical note plus proposed odontogram changes?

- Which X-ray, scanner, and camera devices are realistically available to clinics in Pakistan, and how do they connect?

Out of scope: DICOM PACS integration with hospital-grade CBCT networks, and full radiology reporting — both are natural extensions once the core USB-capture pipeline below is in place.

# 3. End-to-End Workflow at a Glance

The same four-stage pipeline applies whether the source is a still X-ray, a photo, or a short video clip:

- Capture — the device (sensor, camera, or scanner) produces an image or video on the clinic PC via its USB connection.

- Ingest — Dentia's capture layer detects the new image/video and uploads it straight into the patient's Imaging & X-Rays record, tagged with patient ID, tooth number (if selected), and modality.

- AI Observation — the image (or extracted video frame) is sent to Groq's vision model with a dental-specific prompt; Groq returns structured findings in JSON plus a narrative note.

- Chart & Note Integration — the findings populate a draft AI clinical note and a draft set of odontogram changes; the dentist reviews both and clicks Save to commit them — nothing writes to the permanent record automatically.

Manual Save Safety applies here exactly as it does in the Ortho/TMJ diagnostic suite: AI-proposed chart changes are staged, never auto-committed. The dentist always has final sign-off.

# 4. Imaging & Video Device Categories

Chairside capture hardware falls into four categories, each with a different connection method and a different route into Dentia.

## 4.1  Intraoral X-Ray Sensors (RVG / Digital Radiography)

A radiovisiography (RVG) sensor is the small cabled chip placed against a tooth in place of film. When the X-ray head fires, the sensor converts the exposure into a digital periapical or bitewing image in one to two seconds, at a significant radiation-dose reduction versus film. Sensors connect over a standard USB cable and are treated by Windows as a TWAIN or vendor-SDK imaging source, not a plug-and-play webcam — capture is triggered by the sensor's own driver software, not a generic browser camera call.

- Typical use: periapicals, bitewings, working-length checks during root canal, caries and bone-level assessment, implant planning alongside CBCT.

- Connection: USB 2.0 cable, proprietary TWAIN/SDK driver bundled by the sensor manufacturer.

## 4.2  Panoramic & Cephalometric (OPG) Machines

Panoramic (orthopantomogram / OPG) and cephalometric units are wall-mounted or floor-standing X-ray machines that capture a full-jaw or skull-profile image in a single rotating exposure. These are not USB peripherals of the clinic PC in the same sense as a sensor — the machine has its own imaging computer, and the resulting image is exported to the practice-management workstation either over the clinic network (DICOM) or via a TWAIN bridge into imaging software, then dropped into a shared folder or PACS location.

- Typical use: full-arch overview, wisdom-tooth and impaction assessment, orthodontic and TMJ evaluation, implant planning.

- Connection: network DICOM export or TWAIN-bridged software integration (not raw USB webcam capture).

## 4.3  Intraoral Cameras (Chairside Photo/Video)

An intraoral camera is a pen-shaped wand with an LED-lit lens that a dentist moves around the mouth to show a patient — and to document — a tooth or soft-tissue area up close. Unlike an X-ray sensor, most intraoral cameras are USB Video Class (UVC) devices, meaning the operating system and browser both recognize them exactly like a standard webcam: no special driver is required, and a photo or live video stream can be pulled directly through the browser's camera API.

- Typical use: patient education, cavity/gumline documentation, before-after photos, cosmetic case presentation.

- Connection: USB (UVC-class, plug-and-play) or Wi-Fi/wireless models that stream to a companion app.

## 4.4  Webcam / Doctor-Facing Video (Tele-Consult & Chairside Notes)

A standard USB or built-in laptop webcam can record short clips during a live consultation — for example, a dentist narrating findings while pointing at a tooth, or a quick video of jaw movement for a TMJ assessment. Like intraoral cameras, this is UVC-class hardware and is captured the same way through the browser, with the resulting clip (or a handful of extracted frames) sent onward to the AI pipeline exactly like a photo.

# 5. How Each Device Connects to Dentia

The critical distinction for the engineering team is that browsers can talk directly to UVC cameras, but cannot talk directly to X-ray sensors or OPG machines — those require a small local companion application. The table below maps each device type to its actual connection path.

| Device Type | Connects As | Capture Method Into Dentia |
| --- | --- | --- |
| Intraoral X-ray sensor (RVG) | TWAIN / vendor SDK imaging source | Local Capture Bridge watches the sensor's export folder or calls its SDK, then auto-uploads to Imaging & X-Rays |
| Panoramic / Cephalometric (OPG) | Network DICOM export or TWAIN bridge | Local Capture Bridge polls the shared export path / DICOM node and auto-uploads new studies |
| Intraoral camera (wired) | USB Video Class (UVC) — like a webcam | Browser captures directly via getUserMedia(); “Capture” button saves the frame straight to the patient record |
| Intraoral camera (Wi-Fi/wireless) | Vendor app streaming to browser or local network | Companion app relays the feed to a local port the browser reads, or vendor SDK pushes frames to Dentia's upload endpoint |
| Webcam (consultation video) | USB Video Class (UVC) | Browser records short clips via MediaRecorder API; clip or key frames upload on Stop |

## 5.1  The Local Capture Bridge (for X-ray & OPG devices)

Because browsers are sandboxed away from TWAIN drivers and DICOM ports for security reasons, X-ray sensors and OPG machines need a small always-on Windows helper — the Capture Bridge — installed once on the clinic PC. It does three things:

- Watches the folder (or DICOM node) the sensor/OPG software writes new images to.

- On a new file, tags it with the active patient ID and appointment from Dentia's open session, then uploads it to the Imaging & X-Rays module over HTTPS.

- Notifies the open Dentia browser tab in real time (via a local WebSocket) so the new X-ray appears on screen within a second or two of exposure — no manual “browse and upload” step.

This mirrors how existing practice-management systems (Dentrix, Eaglesoft, Open Dental) already integrate Schick, Dexis, Carestream, and Vatech sensors — a lightweight bridge service, not a browser plugin, is the standard pattern for TWAIN-based dental hardware.

## 5.2  Browser-Native Capture (for intraoral cameras & webcams)

UVC cameras need no bridge. Dentia's web app requests camera access the same way a video-call app does, lists the connected camera(s) — including a USB intraoral camera appearing alongside any built-in webcam — lets the dentist pick the device and preview the live feed, and offers a Capture Photo or Record Clip control right inside the patient chart. The saved photo or clip is uploaded to Imaging & X-Rays immediately, tagged to the current patient, tooth (if selected on the odontogram first), and visit.

# 6. Groq AI Observation Pipeline

Once an image (or a small set of extracted video frames) lands in Imaging & X-Rays, Dentia's backend sends it to Groq's multimodal API for analysis. Groq is used here specifically for its very low inference latency — the goal is a clinical read appearing in seconds while the dentist is still chairside, not a batch job reviewed later.

## 6.1  Model & API Call

Groq's current vision-capable model (a 27B-parameter multimodal model supporting image plus text input and native JSON-mode output) receives the image alongside a dental-specific system prompt over Groq's OpenAI-compatible chat completions endpoint. The image is base64-encoded and embedded in the request; Groq returns a response in well under a second for a single image.

- Endpoint: chat completions API (OpenAI SDK-compatible), model set to Groq's current vision model.

- Input: base64 image + a system prompt constraining the model to dental findings and Dentia's JSON schema.

- Output mode: JSON mode, so the response is guaranteed parseable — no free-text scraping.

## 6.2  Prompt Design

The system prompt does three jobs: it tells the model what kind of image it is looking at (periapical X-ray, panoramic, intraoral photo, or video frame), constrains its vocabulary to Dentia's existing condition and CDT taxonomy (Section 3 of the Clinical Manual), and forces tooth-number references to the active numbering system (Universal or FDI) already selected for that patient.

- Modality context: “This is a periapical radiograph of tooth #19” or “This is a live intraoral photo, unlabeled — identify the visible tooth/teeth.”

- Constrained vocabulary: caries, periapical radiolucency, bone loss, fracture, calculus, restoration, missing tooth, implant, impaction — matched to Dentia's condition legend and CDT code list.

- Confidence flag: the model scores its own certainty per finding so low-confidence reads are visually flagged for the dentist rather than silently applied.

## 6.3  Structured Output Schema

Groq's JSON mode returns a fixed schema so the backend can act on it programmatically rather than parsing prose:

| Field | Example Value | Purpose |
| --- | --- | --- |
| toothNumber | "19" | Maps the finding to a specific odontogram tooth |
| numberingSystem | "Universal" | Confirms Universal vs FDI so the chart maps correctly |
| surface | ["D", "O"] | Which of the 5 surfaces (MODBL) are affected |
| finding | "Periapical radiolucency, ~3mm, distal apex" | Human-readable clinical observation |
| suggestedCondition | "Root Canal Indicated" | Maps to Dentia's condition color legend |
| suggestedCdtCode | "D3330" | Draft CDT code for the proposed procedure |
| confidence | 0.86 | Model's self-reported certainty (0–1) |
| sourceImageId | "radiograph_4471" | Links the finding back to the original image in Imaging & X-Rays |

# 7. AI Clinical Notes Generation

The structured findings from Section 6.3 are fed into the same 8-section SOAP template used by the Ambient AI Voice Scribe (Section 5.7 of the Clinical Manual), so an X-ray-triggered note and a voice-dictated note share one format and land in the same [dentist].[ClinicalLogs] / [dentist].[DentalNotes] tables.

- Objective Exam section is auto-populated with the imaging findings (e.g. “Periapical radiolucency noted at distal apex of #19 on periapical radiograph”).

- Assessment & Diagnosis section receives a draft ICD-10 / clinical impression derived from the finding.

- Procedures Performed / Recommended is pre-filled with the suggested CDT code, awaiting the dentist's confirmation or edit.

- The source image itself is linked inline in the note, so anyone reviewing the chart later can open the exact X-ray or photo the AI read.

The AI note is always generated as a draft. The dentist reviews, edits if needed, and electronically signs it — identical to the existing voice-scribe workflow — before it becomes part of the permanent record.

# 8. Automated Chart Updates & Safety Controls

The proposed odontogram change (tooth, surface, condition, CDT code) from Section 6.3 is rendered on the chart as a highlighted, pending overlay — visually distinct from confirmed chart data — rather than being written into [dentist].[TeethState] immediately.

- AI finding arrives and is shown as a dashed/pending marker on the relevant tooth, with the source image, finding text, and confidence score visible on hover or tap.

- The dentist taps the marker to accept, edit (e.g. change the surface or condition), or dismiss it.

- Only on Accept does the change get written to [dentist].[TeethState] and reflected in the permanent odontogram — following the same Manual Save Safety principle used throughout Dentia's diagnostic suites.

- Low-confidence findings (below a configurable threshold, e.g. 0.6) are shown but not auto-suggested for one-tap accept — they require the dentist to open the tooth detail view first.

# 9. Imaging & Video Devices Available in Pakistan

The devices below are the categories of hardware realistically sourced by dental clinics in Pakistan today, either through local importers/distributors or regional suppliers serving South Asia. Exact pricing varies by importer and should be confirmed locally; the table focuses on connectivity, since that determines which capture path (Section 5) applies.

## 9.1  Intraoral X-Ray Sensors (RVG)

| Brand / Model Family | Sensor Type | Connectivity | Notes |
| --- | --- | --- | --- |
| NanoPix (Eighteeth) | CMOS, size 1 & 2 | USB 2.0, TWAIN driver | Actively imported and distributed in Pakistan; common entry/mid-tier choice |
| Woodpecker i-Sensor | CMOS | USB 2.0, TWAIN driver | Widely available across South Asia through dental supply importers |
| Carestream / Kodak RVG 5200 | CCD, high resolution | USB, CS Imaging software | Premium tier; strong image quality, higher import cost |
| Vatech EzSensor | CMOS | USB, EzDent-i / TWAIN | Often bundled with Vatech panoramic units for one-vendor workflow |
| Waldent RTA / Siger | CMOS | USB 2.0, TWAIN driver | Budget-friendly, common in smaller clinics |

## 9.2  Panoramic / Cephalometric (OPG) Machines

| Brand / Model Family | Imaging Type | Connectivity | Notes |
| --- | --- | --- | --- |
| Vatech PaX-i series | Panoramic (+ optional Ceph/CBCT) | Network DICOM / TWAIN bridge | Common globally and in the region; upgrade path to CBCT |
| Planmeca ProMax 2D | Panoramic, extraoral bitewing, TMJ | Network DICOM, Romexis software | Premium multi-specialty clinics and radiology centers |
| Carestream CS-series OPG | Panoramic / Ceph | Network DICOM, CS Imaging | Frequently paired with Carestream RVG sensors |

## 9.3  Intraoral Cameras

| Type | Examples in Market | Connectivity | Notes |
| --- | --- | --- | --- |
| Wired USB intraoral camera | Generic UVC pen-cameras, Waldent, MD-series | USB 2.0 (UVC — plug-and-play) | Simplest integration path: works directly with browser capture, no driver needed |
| Wireless/Wi-Fi intraoral camera | HD Wi-Fi endoscope-style cameras sold locally | Wi-Fi to companion app / built-in screen | Needs vendor app or local relay; slightly more setup than wired UVC |

## 9.4  Standard Webcams

Any UVC-compliant USB webcam (Logitech, generic HD webcams widely sold in Pakistan's electronics markets) works out of the box for chairside consultation video — no dental-specific hardware is required for this category.

# 10. Security, Compliance & Data Handling

- Radiation-emitting equipment (X-ray sensors, OPG units) falls under Pakistan Nuclear Regulatory Authority (PNRA) licensing for the physical device and its operation — this is separate from and unaffected by the software integration described here.

- Patient images and AI-derived notes are health data; access should be role-restricted (per the Doctor & Staff Admin module, Section 5.12 of the Clinical Manual) and images should never be sent to Groq or any AI provider without the patient/clinic's data-handling consent already covered under the clinic's standard EHR policy.

- Images in transit to Groq's API should be sent over TLS, and the Capture Bridge should authenticate to Dentia's backend with a clinic-scoped token, not a shared credential.

- Retain the original, unmodified image indefinitely alongside any AI finding — the AI's read is an assistive draft, not a replacement for the source radiograph or photo in the legal record.

# 11. Implementation Roadmap

| Phase | Deliverable | Key Dependency |
| --- | --- | --- |
| Phase 1 | Browser-native capture for UVC intraoral cameras & webcams (photo + short clip) into Imaging & X-Rays | getUserMedia / MediaRecorder browser APIs — no new hardware integration |
| Phase 2 | Local Capture Bridge (Windows service) for TWAIN-based RVG sensors, with folder-watch auto-upload | Sensor vendor TWAIN SDK access; one bridge build per common sensor brand in Section 9.1 |
| Phase 3 | Groq vision pipeline: prompt design, JSON schema, confidence thresholds, pending-chart overlay UI | Groq API key; sample annotated X-ray set for prompt tuning |
| Phase 4 | OPG/Panoramic DICOM or TWAIN-bridge ingestion | Network access to the OPG machine's export path or DICOM node |
| Phase 5 | Full AI Notes + Auto-Chart round-trip with dentist accept/edit/dismiss controls | Phases 1–3 complete; SOAP template integration (Section 5.7 of Clinical Manual) |

# 12. Summary

Two capture paths cover essentially all chairside imaging hardware: UVC devices (intraoral cameras, webcams) connect straight through the browser with no extra software, while X-ray sensors and OPG machines need a small local Capture Bridge because their TWAIN/DICOM interfaces aren't reachable from a sandboxed browser. Either way, the resulting image lands in Imaging & X-Rays automatically and is handed to Groq's vision model for a fast, structured clinical read. That read becomes a draft SOAP note and a set of pending odontogram markers — never a silent chart edit — so the dentist stays the final decision-maker while the AI removes the manual transcription and filing work around every X-ray, photo, and video captured during the visit.